## Slide 1: ADC Calibration in LArTPC
Wenqiang Gu, Brookhaven National Laboratory
for the DUNE collaboration
WIRE SIGNAL
Workshop on Calibration and Reconstruction for LArTPCs
December 10th 2018

## Slide 2: Evolution of ADC in LArTPC
Cold ADC is essential for DUNE
BooNE
Warm
Commercial
AD-9222
protoDUNE
Cold
Non-commercial
“domino” P1
SBND
Cold
Commercial
AD7274
DUNE
Cold
Non-commercial
???
?

## Slide 3: How to choose an ADC?
Dynamic range
  300 fC maximum ionized charge, 0.1 fC noise  12-bit ADC
  Preamp gain 4.7 mV/fC  ADC 0 ~ 1.6V
Digitization frequency
  Nyquist theorem: 1 MHz for 2us shaping
  2 MHz oversampling is helpful for improving resolution, sticky code mitigation, etc.
V. Radeka et al. Cold electronics for ‘Giant’ Liquid Argon Time Projection Chambers,
J. Phys. Conf. Ser. 308 (2011) 012021
protoDUNE: 14 mv/fC + 2.0 µs
Frequency spectrum for different shaping time

## Slide 4: Precise ADC determination in protoDUNE
However, given the cold environment in LAr, two problems occur for the precise determination of ADC
ADC Sticky Code
ADC Nonlinearity
2432
= (100110 000000)2
2367
= (100100 111111)2

## Slide 5: ADC readout scheme in protoDUNE
Read/write logic synchronized through five control signals
In cold environment, an instability in bit conversion results in sticky code and nonlinearity

## Slide 6: Procedure of sticky code mitigation
X
X
X
X
X
X
X
Identify sticky code:
ADC % 64 = 0, 1, 63, …
Veto sticky code candidates
Interpolate the ticks with near-by “non-sticky” ticks

## Slide 7: Sticky code interpolation
Linear interpolation between “un-sticky” codes is a good first step
  2 MHz oversampling is helpful for interpolation
However, linear interpolation may not be sufficient for signal region
A correction w.r.t. the electronics response function would be better
Linear interpolation
W.r.t response function

## Slide 8: FFT interpolation w.r.t electronics response
However, some facts makes it difficult for using the response function
  A few percent channel-to-channel variation in response function
  Coupled with ADC nonlinearity (discussed shortly)
Instead, a FFT interpolation is proposed by
  i) Linear interpolation as a base correction
  ii) Once a “sticky” code found in an even-binned tick, apply phase shift to odd-binned ticks to cover even-binned ticks, and vice versa
FT property | Time domain | Frequency domain
Phase shift |  | 
FT property | Time domain | Frequency domain
Phase shift |  | 
Example: red waveform interpolated from blue

## Slide 9: Advantages of such FT interpolation
Only the phase changed, while no changes of the magnitude in the frequency domain
  Still respect the shape of the electronics response function
Sometimes, good code tagged as “sticky”, the FT interpolation presumably minimize the biases
  Balance of efficiency and accuracy for sticky code tagging

## Slide 10: Performance of the ADC mitigation
Before mitigation
After mitigation
Sticky code
Interpolation

## Slide 11: A special case
However, when two adjacent sticky codes happens on the peak region, the mitigation does not work well
Need to improve this special case
  Maybe ignore the base correction from linear interpolation
Zoom-in
Sticky at 3008 = (101111 000000)2
Preamp pulser waveform

## Slide 12: ADC nonlinearity (NL)
Low temperature degrades the electronics and read/write logics
External clock eases NL as well as sticky code
NL is sensitive to clock settings
A precise determination of ADC NL is very important for the extraction of ionized electrons
noise
nonlinearity
Sticky code at 63
Sticky code at 0

## Slide 13
Noise
Nonlinearity
Sticky code at 63
Sticky code at 0

## Slide 14: Difficulties from a bench test to protoDUNE
WIRE SIGNAL
Bench test
CALIBRATION SIGNAL
(KNOWN VOLTAGE)
ProtoDUNE
No direct voltage input!
NL is sensitive to clock settings
(bench test) clock is tuned for each ADC
(protoDUNE) one clock shared by four ADC circuits

## Slide 15: ADC calibration setup
Similar setup as in MicroBooNE electronics calibration
Preamplifier
Gain *4
Shaping *4
ADC nonlinearity
C. Adams et al., Ionization Electron Signal Processing in Single Phase LArTPCs II. Data/Simulation Comparison and Performance in MicroBooNE
arXiv:1804.02583
Pulser
Voltage * 64
0 ~ 1.4V
Waveform measurement

## Slide 16: Calibration scheme independent of the shape of response function
Assume pulse voltage (η) and preamp gain (G) do not change the shape of electronics response
NL distorts the shape differently for high ADC and low ADC
true ADC
with NL
Waveforms match perfectly in shape
ADCmeasure
ADCtrue
× 1/
× 1/
A spread in shapes
Minimize it!
Input NL in MC

## Slide 17: minimization
ADCmeasure
i-th data point
NL correction function
(To be fitted out)
index of the -th iteration
= pulse voltage × preamp gain
(A normalization of charge input)
Time tick for
is the effective response function
(Calculated with the NL function from best fit of previous iteration)
(, , )

## Slide 18: “Best-fit” and from simulation
Given an initial value of NL correction function
After a few iterations, “best-fit” NL and effective response tends to be stable
The spread in significantly shrinks after minimization
True
Iter. #1
Iter. #2
True
Initial
ADCtrue – ADCmeasure
ADCmeasure
Electronics response bias ≈ 9
NL bias?

## Slide 19: MC validation of the degeneracy
Given a same charge input, the waveform predictions are close (<1 ADC) for
  True response and NL
  A “best-fit” effective response and NL
NL bias in “best-fit” is not a problem!
ADCmeasure
Fitted response fitted NL
True response true NL
By Carlos

## Slide 20: Discussion: ADC impact on physics analysis
ADC nonlinearity calibration is necessary for a precise extraction of ionized electrons
Presumably, an important input for any analysis related with energy/charge
  e.g., particle identification, TPC/light matching
Xin’s talk
TPC/Light Matching
Brooke’s talk
Particle Identification

## Slide 21: Summary
Cold electronics is essential for LArTPC experiments
Sticky code mitigation
  An interpolation approach through FT is proposed and studied in protoDUNE
ADC nonlinearity
  A calibration strategy was proposed and preliminarily studied in simulation